In the last round robin report for styrene (iis06C08) some odd data was observed for AMS (alpha- methylstyrene) in sample #658, while the data of sample #659 were quite normal. Also the spread of the data of sample #658 was much larger than the spread of the data of sample #659.

At the time of the finalisation of this report, no explanation was yet found for this fact.

However, upon further investigation, a probably cause was found that may explain for the odd data set of AMS (alpha-methylstyrene) in sample #658. In UOP 913:92 the information is given that alpha-methylstyrene may co-eleute with 1,2-diethylbenzene (see table 2 on page 5 and the example chromatogram on page 7). Using GC/MS and standard addition with 1,2-diethylbenzene showed that in sample #658 indeed a significant amount of 1,2-diethylbenzene was present (see attached chromatograms).

The assumption that some laboratories could not differentiate between 1,2-diethylbenzene and alpha-methylstyrene due to co-elution of these two components does explain for the bimodal data distribution as observed for sample #658 on page 39 in the final report iis06C08.

And this leads to the conclusion that the 20 laboratories that reported the high results (100.9 +/ 4.8 mg/kg) may have reported the sum of 1,2-diethylbenzene and alpha-methylstyrene and only the 7 laboratories that reported low results (79.5 +/- 4.8 mg/kg) may have reported alpha-methylstyrene results as requested.

Due to the fact that in sample #659 a much smaller amount of 1,2-diethylbenzene was present, the above phenomena did not lead to a significant difference between the two groups of results.