Instructions to participants for the interlaboratory study:

Specific Migration on Food Contact Materials iis16P11SM

*	Please confirm sample receipt as soon as you have received and checked the contents of this package via www.kpmd.co.uk/sgs-iis-cts/. Please give date of receipt according to prescribed date format and press the ”save date” button. Please inform the Institute for Interlaboratory Studies (iis) immediately when something is wrong with the package and/or samples via iisnl@sgs.com.

*	This interlaboratory study concerns a plate of poly propylene (PP) polymer that is to be used in food industry. Each participating laboratory receives one PP plate labelled #16620 for the determination of the specific migration of some heavy metals by total immersion. Please note that the material of the plate is in practice for multi use but for this PT we ask for the first migration step (thus single use option) only. You can check which metals could be determined at the test scope via www.kpmd.co.uk/sgs-iis-cts/.

*	The type of unit for specific migration may be dependent on local regulations. Therefore, it is possible to report the specific migration in different units; mg/dm2, mg/kg and mg/L. Please note that it is not mandatory to report the test results in all these units. However, we need to know the measured “Final concentration in mg/L in simulant” to be able to convert the test results of all participants into the same unit.

*	Please take care to use the following fixed test conditions during this PT.

| Sample #16620 | One poly propylene plate containing some heavy metals |
| --- | --- |
| Simulant | 3% M/V Acetic Acid |
| Exposure time | 2 hrs. |
| Exposure temperature | 100.0°C |
| Migration method | Total immersion |
| Article use | Single use |
| Simulant volume | as per method used (please report volume in additional questions) |

*	Please treat the sample as it was a routine sample and analyse it in the way you normally do in day-to-day circumstances with in mind the conditions mentioned in table above.

*	Reporting of test results

Within a time frame of five weeks test results can be entered or revised. Please report your test results

via www.kpmd.co.uk/sgs-iis-cts/. It is not necessary to enter all test results in one session. Within the time frame you can add (or delete) test results and/or change test results or units. Please do not forget to hit the “Save/Submit” button after entering the test results. It is strongly recommended to login again and to verify if all your test results are present.

*	Please report the analytical results using the indicated units. It is possible to report in the last column the analytical results rounded in accordance with the standard method that was used and in the first result column the same analytical results but less rounded. Please note that it is not mandatory to report both ‘rounded’ and ‘unrounded’ test results. The ‘unrounded’ test results are preferably used for our statistical evaluations. However the ‘rounded’ test results will be used in case the ‘unrounded’ test results are not reported. We suggest to report one extra significant figure in order to give more meaningful statistical calculations. For example, when you use EN13130 for the determination of Specific Migration and you found a result of 0.614 mg Co/dm2, we request you to report 0.6 mg Co/dm2 in accordance with EN13130 and 0.61 mg Co/dm2 as ‘unrounded’ result, but reporting 0.614 mg Co/dm2 will also be OK.

* 	Please enter also the requested test details under ‘Additional Questions’. This might help us in the evaluation of the test results.

*	The official closing date for reporting test results for this PT is October 14, 2016.

*	After the official closing date it is no longer possible to enter or correct test results via the data entry website www.kpmd.co.uk/sgs-iis-cts/. When you cannot report via the website or for unforeseen reasons cannot report before the closing date, please inform the Institute for Interlaboratory Studies (iis).

*	For all communication (e.g. problems with the package/samples, login details for the data entry portal, not reporting in time) or any remarks/questions please contact:

ing. R.J. Starink / Institute for Interlaboratory Studies

P.O. box 200, NL-3200 AE Spijkenisse, The Netherlands

tel.no. +31 181 69 45 41/ fax. +31 181 69 45 43 /	e-mail: iisnl@sgs.com / website: www.iisnl.com